Title:	Running Maddy and Gauge

Document type:

Document ID:	EOI

Author:

Owner:

| Approver(s): |  |
| --- | --- |
| Approval date: |  |
| Effective date: |  |
| Next periodic review date: |  |

# Safety

  - The Accretech UF200R contains moving parts. Be careful of motion and pinch points when working with the chuck, head stage and transfer robot.

Figure 6a: Mechanical Hazards

  - In the event of an emergency, press the red EMO button to stop motion and power off the tool.

Figure 6b: Emergency Off

  - The cassette loader uses a Class 1 laser. Never tamper with the cassette loader lock.

# Procedure

## Setup

    - Verify Accretech UF200R is powered on without warnings (Figure 8a).

### Figure 8a: PROBE08 Powered ON

    - Verify KEYSIGHT 34461A is powered on (Figure 8b).

### Figure 8b: KEYSIGHT 34461A DMM Powered ON

Atomica Corp.	EOI-

Running Maddy and Gauge	Version: 1.0

    - Verify KEITHLEY 2400 is powered on (Figure 8c).

### Figure 8c: KEITHLEY 2400 SMU Powered ON

    - You can also check all instruments by opening the GUI and verifying the green checkmarks on the left.

Figure 8e: Instrument Connections

## Changing Probe Card

Make sure probe card is MADDY_TL_REV D. If so, skip this step.

    - Within the main production screen, press “CARD CHANGE” (Figure 8e).

### Figure 8e: CARD CHANGE

    - Press “CARD CHANGE START” (Figure 8f).

### Figure 8f: CARD CHANGE START

    - Loosen (4) thumb screws and open clamps (Figure 8g).

### Figure 8g: Loosen Screws and Open Clamps

    - Use the hole at the front edge of the probe card to carefully pull the probe card out of the connector socket and remove the probe card (Figure 8h).

WARNING: Be very careful to avoid damaging the probe needles on the bottom of the probe card.

### Figure 8h: Slide Probe Card Out of Edge Connector Socket

    - On the bottom of the probe card, carefully place the plastic probe needle cover over the probe needles and make sure it is secured with tape (Figure 8i).

### Figure 8i: Probe Needle Cover Centered Over Probe Needles

    - Store the old probe card in its case from Accuprobe on the PROBE08 rack (Figure 8j).

### Figure 8j: Store Probe Cards on PROBE08 Rack

    - Carefully remove the plastic probe needle cover from the new probe card and mount it onto the stage rails so that the probe needles are pointing down. Once the new probe card is flat and centered on the rails, use the hole at the front edge of the probe card to push the probe edge connector into the socket (Figure 8k).

### Figure 8k: Use Rails to Guide Probe Card Into Edge Connector Socket

    - Bring the clamps down and tighten (4) thumb screws (Figure 8l).

### Figure 8l: New Probe Card Installed

    - Press “CARD CHANGE END” (Figure 8m).

### Figure 8m: CARD CHANGE END

    - Press “END” (Figure 8n).

### Figure 8n: END in Manual Card Change Prompt

    - Press “END” (Figure 8o) to complete the probe card change and return to the main production screen.

### Figure 8o: END to Complete Card Change

## Changing Device Parameters

    - Next, the correct recipe for the new probe card must be loaded. If the device data is MADDY-STD-SHORT, you may skip this step.

### Figure 8p: DEVICE PARAMETER CHANGE

    - Press “DEVICE LIST REFERENCE” (Figure 8q).

### Figure 8q: DEVICE LIST REFERENCE

    - Verify “HD” is highlighted and press “EXECUTE” (Figure 8r).

### Figure 8r: EXECUTE

    - Select the correct recipe: MADDY-STD-SHORT (Figure 8s).

Figure 8s: Highlight the Intended Recipe in the Device List

    - Press “DEVICE DATA READ” (Figure 8t).

### Figure 8t: DEVICE DATA READ

    - When prompted, press “EXECUTE” (Figure 8u).

### Figure 8u: EXECUTE

    - Press “SETTING END” (Figure 8v) to return to the main production screen.

### Figure 8v: SETTING END

## Starting Probe Routine

    - Press “OPERATION SETTINGS” (Figure 8w).

### Figure 8w: OPERATION SETTINGS

    - Verify the correct Loading/Unloading position is highlighted and press “ENT” (Figure 8x).

### Figure 8x: Loading/Unloading Positions

    - Press “START” (Figure 8y).

### Figure 8y: START

    - Press “Clear screen” when prompted with “Put wafer on CHUCK !!” (Figure 8z).

### Figure 8z: Clear Screen

    - Open the Front Access Door (Figure 8aa).

### Figure 8aa: Opening Front Access Door

    - Load the wafer on top of the chuck pins so that the notch is aligned with the hole at 6 o'clock (Figure 8ab).

### Figure 8ab: Notch Aligned with Hole at 6 O'clock

    - Close the Front Access Door.

    - Press “Start” (Figure 8ac).

### Figure 8ac: START

    - For the first wafer, you may encounter needle alignment error.  If wafer loads without this error, skip past the pin alignment step to the start run section. Press “Buzzer Release” when prompted with “NEEDLE ALIGNMENT ERROR” (Figure 8ad).

### Figure 8ad: Buzzer Release

    - Press “Clear screen” when prompted with another “NEEDLE ALIGNMENT ERROR” (Figure 8ae).

### Figure 8ae: Clear Screen

    - Press “Clear screen” when prompted with “Set CROSS MARK on needle position and push <DATA IN>” (Figure 8af).

### Figure 8af: Clear Screen

    - Press the directional arrows to jog the chuck so that the tips of the probe needles are aligned with the circles overlaying the camera feed (Figure 8ag).

NOTE: The dark green buttons can be used for fine adjustment.

Figure 8ag: Jog Chuck Until Needles Align with Circles

    - Press “DATA IN” (Figure 8ah).

### Figure 8ah: DATA IN

    - Press “Clear screen” when prompted with “STOP AFTER AUTO NEEDLE HEIGHT SETTING DONE” (Figure 8ai).

### Figure 8ai: Clear Screen

    - Verify the CARD RELATIVE Z HEIGHT and Probing Overdrive (Figure 8aj).

NOTE: These are program specific settings. Contact Test/Metrology Engineer for correct settings.

### Figure 8aj: Verify Probe Settings

    - Press “START” (Figure 8ak).

### Figure 8ak: START

    - “Probing” will be displayed in the bottom left corner of the screen (Figure 8al).

### Figure 8al: Ready for Probing

    - Starting Run

- Move to the PROBE08 desktop PC and launch the Atomica Prober executable. (Figure 8am).

NOTE: Login to the PC with the following credentials (User: Atomica, Password: Engineer2022)

### Figure 8am: AtomicaProber.exe/OTTO

- Verify instruments are connected, verify working directory is correct. Load the Maddy ATA folder if not already loaded.

### Figure 8an: Load ATA folder

- Go to recipe tab, check that correct probe card is selected. Check that the recipe is correct (Maddy TL for production, Maddy Gauge for gauge). (Figure 8ao).

### Figure 8ao: Recipe tab, Match the probe card with the label on the physical card.

How to Start Run

Go to the Run tab, make sure recipe is loaded there as well. Press RUN to start.

Note: The prober will not run and throw an error unless it says “Probing” on the bottom left. (Figure 8ap).

### Figure 8ap: Press Run

    - Probing commences.

    - Once probing has completed, press “UNLOAD” under the “Run” tab (Figure 8aq).

### Figure 8aq: UNLOAD

    - Open the Front Access Door and remove the wafer from the chuck, once screen prompts you to.

    - Close the Front Access Door and press “START” on the Accretech UF200R interface (Figure 8ar).

### Figure 8ar: START for Removing Wafer

    - Press “Clear screen" when prompted with ”DONE!!" (Figure 8as).

### Figure 8as: DONE!!

Exporting

Once run is finished, if you are not auto exporting with cassette automation (ie. only running one wafer), go to results tab.

Fill in the correct wafer id and lot id. Make sure export path is as shown (or put it in anywhere else and move it manually). Make sure export format is MaddyTL. Click Export once ready (and save to csv to create the backup results file).

ADD log

Make sure that files are saved in the export folder. Recommended is to have both the results csv and the sql file for each run. The results csv (eg.MADC-V-069_04A7_results.csv) is backup and can be imported back into the gui at any time to see the data and pass fails. (Refer to full manual for steps on this).

The export file (eg. MADC-V-069_03F7_TEST.csv) will need to be uploaded to SPC.

ADD

# Qualiﬁcation Procedure

  - For gage monitor, Load gage wafer MAD-X-008-13B3 using the front access door per Section the Starting Probe Routine section above.

  - Run Maddy TL Gauge recipe instead of Maddy TL.

### Figure 9a: When Ready to Probe, run the gauge recipe

  - In the results tab, enter the LOT ID (MAD-X-008) and WAFER ID (13B3).

  - Press “RUN” to start probing. (Figure 9b).

  - Once probing is complete, unload the wafer per the Starting Probe Routine steps above.

  - Go to results tab and export. Verify the data file saved at default file path \\prober\ETL\RAWDATA\PROBE08 (Figure 9c).

### Figure 9c: Gage Wafer Data Files on Prober

  - Save a copy of the data file to C:\Data on a lab PC. Rename the file as “savedatahere.csv”.

  - Open SPC Chart CC-02329 PROBE08 (navigate to IMT Equip/Proc/IQA Mon => GAGES => CC-02329 PROBE08). Press the “Add Data” button to upload data from the data file in C:\Data. Verify the data then press “Save”.

  - Use data to verify the process is in control. Put PROBE08 down in MES if data is out of control limits.